## threads
modt_id
HEATER
MAINPUMP
VALVE
SENSORS
BMS
LOGGER
PRESSURE
UI_COMM
UC_MASTER
Note on A2: Sikkandar Sulaiman Ayubkhan: Thread details are given in sheets with name <modt_id>_msg_id. For example, this file contains heater_msg_id sheet.

## cmds
thread_cmd
INIT | /* thread based commands */
START
RUN
STOP
DEINIT
EXIT
GET_CFG | /* config related commands - msg.param contains config id */
SET_CFG
UNSET_CFG
EXEC | /* generic - msg.param contains module specific command */
STATUS
ASSERT_ERROR

## bms_msg_id
filename | config | tcmd | status
BMS/bms.h | CHARGING_STATUS | START_CHARGING
| BAT_TEMP | STOP_CHARGING
BAT_VOLT
BAT_CURRENT
BAT_AVG_CRNT
RELATIVE_SOC
ABSOLUTE_SOC
REM_CAPACITY
FULL_CHG_CAPACITY
RUN_TIME_TO_EMPTY
AVG_TIME_TO_EMPTY
AVG_TIME_TO_FULL
BAT_CHG_CURRENT
BAT_CHG_VOLT
BAT_DSCHG_COUNT
BAT_CELL1_VOLT
BAT_CELL2_VOLT
BAT_CELL3_VOLT
BAT_HEALTH

## heater_msg_id
filename | config | tcmd | status
heater/heater.h | RUN_INTERVAL
MAX_PWM_DC
MIN_PWM_DC
TARGET_FLUID_TEMP
MAX_HEATER_TEMP
ADC_READ_TIMEOUT
ADC_LPF_WIN_SIZE
INLET_TEMP
OUTLET_TEMP
HPAD1_TEMP
HPAD2_TEMP
PWMOUT
HT_MODE

## mainpump_msg_id
filename | config | tcmd | status
mainpump/mainpump.h | CUR_PWM_DC | INC_PWM_DC
| PWM_INC_DEC_STEP | DEC_PWM_DC
UPDATE_FLOWRATE

## valve_msg_id
filename | config | tcmd | status
valve/valve.h | MOVE_DURATION | OPEN
| CONT_MOVEMENT | CLOSE
PATTERN

## sensors_msg_id
filename | config | tcmd | status
sensor/sensor.h | TURB_TOP_RUN_INTERVAL | ENABLE_ALL
| BUB_RUN_INTERVAL | DISABLE_ALL
| BUB_LPF_WIN_SIZE | ENABLE_BUB
| TURB_TOP_VAL | DISABLE_BUB
| TURB_SIDE_VAL | ENABLE_TURB_TOP
| BUBBLE_DET_VAL | DISABLE_TURB_TOP
| TURB_TOP_GAIN | ENABLE_TURB_SIDE
| TURB_SIDE_GAIN | DISABLE_TURB_SIDE
| TURB_TOP_INT_TIME | ENABLE_PERITONITIS
| TURB_SIDE_INT_TIME | DISABLE_PERITONITIS
| DOOR_STATUS | ENABLE_CURRENT_H1
| CASSETTE_STATUS | DISABLE_CURRENT_H1
| HALL_1_STATUS | ENABLE_CURRENT_H2
| PERIT_R_VAL | DISABLE_CURRENT_H2
PERIT_G_VAL
PERIT_B_VAL
CURRENT_H1_VAL
CURRENT_H2_VAL

## logger_msg_id
filename | config | tcmd | status
logger/logger.h | EVENT_LOG | LOG_TEST_POINT
DATA_LOG

## pressure_msg_id
filename | config | tcmd | status
pressure/pressure_monitor.h | RUN_INTERVAL | ZERO_OFFSET
| LC1_ACT_RAW_VAL | 2PT_CALIB
| LC2_ACT_RAW_VAL | ENABLE_MONITOR
| LC1_REL_RAW_VAL | DISABLE_MONITOR
| LC2_REL_RAW_VAL | AIRPUMP_ON
| LC1_REL_PA_VAL | AIRPUMP_OFF
LC2_REL_PA_VAL
AIR_ACT_PA_VAL
AIR_REL_PA_VAL

## ui_comm_msg_id
filename | config | tcmd | status
communication/ui_comm.h |  | UI_RX_MSG
UI_TX_MSG
TX_LOG_CFG_1T
TX_LOG_CFG_ALL
RX_FILE_CHECK
RX_FILE_CHUNK
RX_FILE_ABORT
UART_RX_ERR
TS_HANDSHAKE_1

## uc_master_msg_id
filename | config | tcmd | status
usecase/uc_master.h | CCPD_POST_BEGIN_MS | UC_DEMO_START | CCPD_POST
| CCPD_POST_MAINPUMP | UC_CCPD_START | CCPD_SYSCHK
| CCPD_POST_BATTERY | UC_RUN_POST | CCPD_CALIB
| CCPD_POST_HT_INLET | UC_RUN_SYSCHK | CCPD_PRIME
| CCPD_POST_HT_OUTLET | UC_RUN_CALIB | CCPD_FILL
| CCPD_POST_HT_HPAD1 | UC_RUN_PRIME | CCPD_DWELL
| CCPD_POST_HT_HPAD2 | UC_RUN_THERAPY | CCPD_DRAIN
| CCPD_POST_HEATER_HEAT | UC_ABORT
| CCPD_POST_TURB_TOP | UC_NEXT_STEP
| CCPD_POST_TURB_SIDE | UC_NEXT_SUBSTEP
| CCPD_POST_PERIT | UC_RETRY_STEP
| CCPD_POST_LEVEL | UC_RETRY_SUBSTEP
CCPD_POST_AIRPA
CCPD_POST_EEPROM
CCPD_POST_FAN
CCPD_POST_END_MS
CCPD_RUNNING_STATE

## data_log_type
msg_id | ctype | datalog_column_name | datalog_column_unit | enable | dom_display_id
MSG_HEATER_GET_RUN_INTERVAL | uint16 | heater run interval | ms | 0
MSG_HEATER_GET_MAX_PWM_DC | uint8 | heater max pwm dc | % | 1 | disp_heater_max_pwm_dc
MSG_HEATER_GET_MIN_PWM_DC | uint8 | heater min pwm dc | % | 0
MSG_HEATER_GET_TARGET_FLUID_TEMP | int16 | target fluid temp | degC | 1
MSG_HEATER_GET_MAX_HEATER_TEMP | int16 | max heater temp | degC | 1 | disp_heater_max_hpad_t1
MSG_HEATER_GET_ADC_LPF_WIN_SIZE | uint8 | heater adc lpf size | - | 0
MSG_HEATER_GET_INLET_TEMP | int16 | inlet fluid temp | degC | 1 | disp_inlet_temp
MSG_HEATER_GET_OUTLET_TEMP | int16 | outlet fluid temp | degC | 1 | disp_outlet_temp
MSG_HEATER_GET_HPAD1_TEMP | int16 | hpad1 temp | degC | 1 | disp_hpad_1_temp
MSG_HEATER_GET_HPAD2_TEMP | int16 | hpad2 temp | degC | 1 | disp_hpad_2_temp
MSG_HEATER_GET_PWMOUT | uint8 | heater pwm dc | % | 1 | disp_heater_pwm_dc
MSG_MAINPUMP_GET_CUR_PWM_DC | uint8 | mainpump pwm dc | % | 1 | disp_mainpump_pwm_dc
MSG_VALVE_GET_MOVE_DURATION | uint16 | valve move time | ms | 0
MSG_SENSORS_GET_TURB_TOP_RUN_INTERVAL | uint16 | turbidity top sample time | ms | 0
MSG_SENSORS_GET_TURB_SIDE_RUN_INTERVAL | uint16 | turbidity side sample time | ms | 0
MSG_SENSORS_GET_BUB_LPF_WIN_SIZE | uint8 | bubble lpf size | - | 0
MSG_SENSORS_GET_BUB_RUN_INTERVAL | uint16 | bubble sample time | ms | 0
MSG_SENSORS_GET_TURB_TOP_VAL | uint16 | turb top value | ir | 1 | disp_turb_top
MSG_SENSORS_GET_TURB_SIDE_VAL | uint16 | turb side value | ir | 1 | disp_turb_side
MSG_SENSORS_GET_BUBBLE_DET_VAL | uint8 | level status | - | 1 | disp_level_st
MSG_SENSORS_GET_DOOR_STATUS | uint8 | door opened | - | 1 | disp_door_st
MSG_SENSORS_GET_CASSETTE_STATUS | uint8 | cassette inserted | - | 1 | disp_cassette_st
MSG_SENSORS_GET_HALL_1_STATUS | uint8 | hall status | - | 1 | disp_hall_st
MSG_SENSORS_GET_CURRENT_H1_VAL | uint32 | hpad1 current | mA | 1 | disp_curr_h1
MSG_SENSORS_GET_CURRENT_H2_VAL | uint32 | hpad2 current | mA | 1 | disp_curr_h2
MSG_PRESSURE_GET_LC1_REL_PA_VAL | float32 | load cell 1 pressure | mBar | 1 | disp_lc1_pa
MSG_PRESSURE_GET_LC2_REL_PA_VAL | float32 | load cell 2 pressure | mBar | 1 | disp_lc2_pa
MSG_PRESSURE_GET_AIR_ACT_PA_VAL | float32 | air pressure act | mBar | 1 | disp_airpa_val_act
MSG_PRESSURE_GET_AIR_REL_PA_VAL | float32 | air pressure rel | mBar | 1 | disp_airpa_val_rel
MSG_PRESSURE_GET_LC1_ACT_RAW_VAL | int32 | load cell 1 raw act | raw | 1 | disp_lc1_act
MSG_PRESSURE_GET_LC2_ACT_RAW_VAL | int32 | load cell 2 raw act | raw | 1 | disp_lc2_act
MSG_PRESSURE_GET_LC1_REL_RAW_VAL | int32 | load cell 1 raw rel | raw | 1 | disp_lc1_rel
MSG_PRESSURE_GET_LC2_REL_RAW_VAL | int32 | load cell 2 raw rel | raw | 1 | disp_lc2_rel
MSG_MAINPUMP_GET_CUR_PWM_DC | uint8 | mainpump pwm dc | % | 1 | disp_mainpump_pwm_dc
MSG_SENSORS_GET_PERIT_R_VAL | uint8 | peritonitis red |  | 1 | disp_perit_red
MSG_SENSORS_GET_PERIT_G_VAL | uint8 | peritonitis green |  | 1 | disp_perit_green
MSG_SENSORS_GET_PERIT_B_VAL | uint8 | peritonitis blue |  | 1 | disp_perit_blue
MSG_BMS_GET_CHARGING_STATUS | uint8 | battery charging |  | 1 | disp_bat_chg_en
MSG_BMS_GET_BAT_TEMP | float32 | battery temperature | degC | 1 | disp_bat_temp
MSG_BMS_GET_BAT_VOLT | float32 | battery voltage | V | 1 | disp_bat_volt
MSG_BMS_GET_BAT_CURRENT | int16 | battery current | mA | 1 | disp_bat_current
MSG_BMS_GET_BAT_AVG_CRNT | int16 | battery average current | mA | 1 | disp_bat_avg_current
MSG_BMS_GET_RELATIVE_SOC | uint8 | battery relative soc | % | 1 | disp_rel_soc
MSG_BMS_GET_ABSOLUTE_SOC | uint8 | battery absolute soc | % | 1 | disp_abs_soc
MSG_BMS_GET_REM_CAPACITY | uint16 | battery rem capacity | mA | 1 | disp_rem_capacity
MSG_BMS_GET_RUN_TIME_TO_EMPTY | uint16 | battery run time to empty | minutes | 1 | disp_run_tte
MSG_BMS_GET_AVG_TIME_TO_EMPTY | uint16 | battery avg time to empty | minutes | 1 | disp_avg_tte
MSG_BMS_GET_AVG_TIME_TO_FULL | uint16 | battery avg time to full | minutes | 1 | disp_avg_ttf
MSG_BMS_GET_BAT_CHG_CURRENT | uint16 | battery charging current | mA | 1 | disp_chg_current
MSG_BMS_GET_BAT_CHG_VOLT | float32 | battery charging volt | V | 1 | disp_chg_volt
MSG_BMS_GET_BAT_DSCHG_COUNT | uint8 | battery discharge cycle count |  | 1 | disp_dschg_cycle
MSG_BMS_GET_BAT_CELL1_VOLT | float32 | battery cell 1 volt | V | 1 | disp_cell1_volt
MSG_BMS_GET_BAT_CELL2_VOLT | float32 | battery cell 2 volt | V | 1 | disp_cell2_volt
MSG_BMS_GET_BAT_CELL3_VOLT | float32 | battery cell 3 volt | V | 1 | disp_cell3_volt
MSG_BMS_GET_BAT_HEALTH | uint8 | battery health | % | 1 | disp_bat_health
MSG_UC_MASTER_GET_CCPD_RUNNING_STATE | uint32 | system state |  | 1

## event_log_type
msg_id | enable | expect_timestamp | dom_trigger_id | positive | negative
MSG_HEATER_INIT | 1 | 1 | trig_heater_start | Heater init success | Heater init failure
MSG_HEATER_START | 1 | 1 | trig_heater_start | Heater start success | Heater start failure
MSG_HEATER_STOP | 1 | 1 | trig_heater_stop | Heater stop success | Heater stop failure
MSG_HEATER_DEINIT | 1 | 1 | trig_heater_stop | Heater deinit success | Heater deinit failure
MSG_PRESSURE_INIT | 1 | 1 | trig_pressure_start | Pressure init success | Pressure init failure
MSG_PRESSURE_START | 1 | 1 | trig_pressure_start | Pressure start success | Pressure start failure
MSG_PRESSURE_EXEC_ZERO_OFFSET | 1 | 1 | trig_pressure_zero_off | Pressure zero offset done | Pressure zero offset failed
MSG_PRESSURE_EXEC_2PT_CALIB | 1 | 1 | trig_pressure_2pt_calib | Pressure 2pt calib done | Pressure 2pt calib failed
MSG_PRESSURE_EXEC_AIRPUMP_ON | 1 | 1 | trig_airpump_on | Airpump on | Airpump on failed
MSG_PRESSURE_EXEC_AIRPUMP_OFF | 1 | 1 | trig_airpump_off | Airpump off | Airpump off failed
MSG_PRESSURE_STOP | 1 | 1 | trig_pressure_stop | Pressure stop success | Pressure stop failure
MSG_PRESSURE_DEINIT | 1 | 1 | trig_pressure_stop | Pressure deinit success | Pressure deinit failure
MSG_SENSORS_INIT | 1 | 1 | trig_sensor_start | Sensors init success | Sensors init failed
MSG_SENSORS_START | 1 | 1 | trig_sensor_start | Sensors start success | Sensors start failed
MSG_SENSORS_SET_TURB_TOP_GAIN | 1 | 1 | trig_set_turb_top_gain | Turb top gain set | Turb top gain failed
MSG_SENSORS_SET_TURB_SIDE_GAIN | 1 | 1 | trig_set_turb_side_gain | Turb side gain set | Turb side gain failed
MSG_SENSORS_SET_TURB_TOP_INT_TIME | 1 | 1 | trig_set_turb_top_inttime | Turb top int time set | Turb top int time failed
MSG_SENSORS_SET_TURB_SIDE_INT_TIME | 1 | 1 | trig_set_turb_side_inttime | Turb side int time set | Turb top int time failed
MSG_SENSORS_SET_BUBBLE_DET_VAL | 1 | 1 | trig_set_level_val | Level value overridden | Level set value failed
MSG_SENSORS_UNSET_BUBBLE_DET_VAL | 1 | 1 | trig_unset_level_val | Level value real time | Level real time failed
MSG_UI_COMM_EXEC_TS_HANDSHAKE_1 | 1 | 1 |  | Handshake success | Handshake failure
MSG_HEATER_SET_MAX_PWM_DC | 1 | 1 | trig_heater_max_pwm_dc | Max heater pwm updated | Max heater pwm failed
MSG_HEATER_SET_MAX_HEATER_TEMP | 1 | 1 | trig_heater_max_hpad_t1 | Max heater temp updated | Max heater temp set fail
MSG_MAINPUMP_INIT | 1 | 1 | trig_mainpump_start | Mainpump init success | Mainpump init failed
MSG_MAINPUMP_START | 1 | 1 | trig_mainpump_start | Mainpump start success | Mainpump start failed
MSG_MAINPUMP_SET_CUR_PWM_DC | 1 | 1 | trig_mainpump_set_pwm_dc | Mainpump pwm updated | Mainpump pwm set failed
MSG_MAINPUMP_EXEC_UPDATE_FLOWRATE | 1 | 1 | trig_mainpump_set_flowrate | Mainpump flow updated | Mainpump flow set failed
MSG_MAINPUMP_STOP | 1 | 1 | trig_mainpump_stop | Mainpump stop success | Mainpump stop failed
MSG_MAINPUMP_DEINIT | 1 | 1 | trig_mainpump_deinit | Mainpump deinit success | Mainpump deinit failed
MSG_UC_MASTER_EXEC_UC_CCPD_START | 1 | 1 | trig_uc_ccpd_start | CCPD ready to start | CCPD start failed
MSG_UC_MASTER_EXEC_UC_RUN_POST | 1 | 1 | trig_uc_run_post | POST started | POST start failed
MSG_UC_MASTER_EXEC_UC_RUN_SYSCHK | 1 | 1 | trig_uc_run_syschk | System check started | System check start failed
MSG_UC_MASTER_EXEC_UC_RUN_CALIB | 1 | 1 | trig_uc_run_calib | Calibration started | Calib start failed
MSG_UC_MASTER_EXEC_UC_RUN_PRIME | 1 | 1 | trig_uc_run_prime | Priming started | Prime start failed
MSG_UC_MASTER_EXEC_UC_RUN_THERAPY | 1 | 1 | trig_uc_run_therapy | Therapy started | Therapy start failed
MSG_UC_MASTER_EXEC_UC_ABORT | 1 | 1 | trig_uc_abort | UC abort success | UC abort failure
MSG_UC_MASTER_EXEC_UC_NEXT_STEP | 1 | 1 | trig_uc_next_step
MSG_UC_MASTER_EXEC_UC_NEXT_SUBSTEP | 1 | 1 | trig_uc_next_sub | UC moving next substep | UC next substep failed
MSG_UC_MASTER_EXEC_UC_RETRY_STEP | 1 | 1 | trig_uc_retry_step
MSG_UC_MASTER_EXEC_UC_RETRY_SUBSTEP | 1 | 1 | trig_uc_retry_sub
MSG_BMS_EXEC_START_CHARGING | 1 | 1 | trig_en_bat_chg | started charging battery | battery charging failed
MSG_BMS_EXEC_STOP_CHARGING | 1 | 1 | trig_dis_bat_chg | stopped charging battery | battery charging failed
MSG_BMS_INIT | 1 | 1 | trig_bms_start | BMS init success | BMS init failure
MSG_BMS_START | 1 | 1 | trig_bms_start | BMS start success | BMS start failure
MSG_VALVE_EXEC_PATTERN | 1 | 1 | trig_valve_pattern | valve actuated | valve actuation failure
MSG_VALVE_EXEC_OPEN | 1 | 1 | trig_valve_open | valve actuated | valve actuation failure
MSG_VALVE_EXEC_CLOSE | 1 | 1 | trig_valve_close | valve actuated | valve actuation failure
MSG_VALVE_INIT | 1 | 1 | trig_valve_init | valve init success | valve init failure

## mod_status
MOD_STATUS_POSITIVE | MOD_STATUS_NEGATIVE